CMG-3ESP CALIBRATION SHEET

| WORKS ORDER: | 1637 | DATE: | 15/05/01 |
| --- | --- | --- | --- |
| SERIAL NUMBER: | T3A60 | TESTED BY: | SDG |

|  | Velocity | Mass Position | Feedback Coil |
| --- | --- | --- | --- |
|  | Output | Output | Constant |
|  | V/m/s | (Acceleration | Amp/m/s2 |
|  | (Differential) | output) |  |
|  |  | V/m/s2 |  |
| VERTICAL | 2 x 2946 | 922 | 0.01961 |
| NORTH/SOUTH | 2 x 2966 | 976 | 0.02076 |
| EAST/WEST | 2 x 2968 | 974 | 0.02072 |

| Power Consumption: | 57mA @ +12V input |
| --- | --- |
| Calibration Resistor: | 51000 |

NOTE:  A factor of 2 x must be used when the sensor outputs are used differentially (also known as push-pull or balanced output).  Under no conditions should the negative outputs be connected to the signal ground.  A separate signal ground pin is provided.

POLES AND ZEROS TABLE

WORKS ORDER NUMBER: 1637

SENSOR SERIAL NO: T3A60

Velocity response output, Vertical Sensor:

| POLES (HZ) | ZEROS HZ |
| --- | --- |
| -23.56 x 10-3 ( j 23.56 x 10-3 | 0 |
| -180 | 0 |
| -160 |  |
| -80 |  |
| Normalizing factor at 1 Hz: A = | 2.304 x 106 |

Sensor Sensitivity:		See Calibration Sheet.

Velocity response output, Horizontal Sensors:

| POLES (HZ) | ZEROS (HZ) |
| --- | --- |
| -23.56 x 10-3 ( j 23.56 x 10-3 | 0 |
| -180 | 0 |
| -160 |  |
| -80 |  |
| Normalizing factor at 1 Hz: A = | 2.304 x 106 |

Sensor Sensitivity:		See Calibration Sheet.

NOTE:  The above poles and zeros apply to the vertical and the horizontal sensors and are given in units of Hz.  To convert to Radian/sec multiply each pole or zero with 2  The normalizing factor A should also be recalculated.