## 3  Work of LaCoste under Romberg's direction

As a student of physics under Prof. Romberg at the University of Texas at Austin, LeCoste addressed the problem of how to build a compact, low  effective spring in the 1930's. He inclined the spring at an angle with respect to vertical and provided twist in the coiling operation during fabrication. In another webplace, quantitative reasons are provided for why this zero-length design by LeCoste is a superior one for vertical seismometers. His patent has stood the test of time in terms of improved mechanical stability.

### 3.1  Modern instruments-Force balance technique

Many forms of modern seismometers do not use passive sensing as described above. Instead, a transducer is involved in a force-balance feedback network. In the absence of feedback, acceleration of the case would cause a change in y as indicated above. With the transducer and network, the feedback signal is adjusted to whatever value is necessary for y  =  0. In turn, the error signal (proportional to the amount of feedback) is a measure of how y would otherwise change. Such a system is always equivalent to a harmonic oscillator in its simplest form, and thus to a mass/spring system. Instead of relying totally on the mechanical spring to determine the dynamics of the instrument, the phase and amplitude of the feedback signal is adjusted to yield a seismometer with an effective spring constant eff, whose magnitude can be varied. [Note: W/L describe the force-balance seismometer in the following way: "The idea behind a conventional LP seismometer is to measure the motion of the ground against an inertial reference, i.e., against an elastically suspended mass that is supposed to stay at rest when the ground moves. A force-balance seismometer is a negative-feedback (i.e., electronic servo) system that causes the mass to follow the motion of the ground.  The force required for this purpose is a measure of the ground acceleration.''] It should be noted that eff of the force-balance instrument must be small if the sensitivity is to be large, just as with the conventional instrument. I don’t see why that should be.  To think that modern instruments are sensitive even though they have a short period (when operating without feedback) is to miss the point altogether.  The point being?

3.2  Limitations of the Force-balance approach

An analysis of limitations inherent to the force-balance method requires some careful thought. Ostensibly, the method appears to solve many problems encountered in seismometer design. There are, however, serious issues of anelasticity that have never apparently been considered in relationship to the use of these instruments.

### 3.3  Gedanken to illustrate that limitations exist

Lest one believe that force-balance feedback is infinitely superior to conventional seismometry, consider the following logic. Why even bother with the leaf-spring that is commonly used to support the test mass in these instruments?  I have done a design for an instrument which would not require the use of a spring at all.  It might not be completely practical as a commercial design, but it would detect ground motion as well or better than the present design using a spring.  Why not just add a feedback network to a solid state mass balance instrument that works with resistive strain gauges? Place a big test mass on the pan of the modified mass balance (mmb), add a magnetic transducer of some type to provide a significant lifting force on the mass, and ``voila''-with proper feedback adjustment we suddenly can see earthquakes with the simplest of instruments.  Hopefully everyone will quickly recognize the folly of this reasoning and know that such a modified mass measuring instrument is not capable of functioning as a bonafide seismometer. I see no reason at all why this would not work.  Of course a strain gauge has a much smaller dynamic range than the capacitive sensor we are using, so it would not make a very good instrument, but it should detect the larger ground motions quite adequately.  But why? The answer to this question lies in the following observation. System adaptability [??] is no better than the integrity of the ``spring'' used in generation of the error signal. As noted earlier, any error signal requires the measurement of strain. It, in fact, requires the measurement of the mass position—no strain required.  In the case of the hypothetical modified mmb, the ``spring'', in the absence of feedback, has an exceedingly large . In the case of the W/S leaf-spring seismometer, the leaf has a considerably larger  than that of the conventional seismometer.  Can electronics soften even the hardest springs?  Viewed correctly, feedback should rather be thought of as ‘hardening’ the spring.  Place a coin on the boom, and feedback will quickly have it restored to a point which is within relatively few nanometers of its initial level.  The instrument’s ‘‘ (change in position)(change in force) is extremely small.  The answer is obviously no! What are the limitations to softening? I submit to the reader that there are a host of unanswered questions in the matter. I would submit that feedback seismometer designs are quite well understood by the people who design them.  It is easy to see that electronics limitations (addressed earlier) pose an ultimate upper limit on the size of .

But anelasticity of the support is probably even more important than the electronics-and the problems borne of it are mostly unstudied. This is true in spite of the fact that practictioners understand that an instrument must be allowed to settle for some time after initial loading, before it becomes dependable. This settling is necessary to minimize the effects of anelasticity, through a type of work-hardening.  I take issue with the use of the term ‘work-hardening’ here.  To me, that is an effect associated with bending, forging, rolling, drawing, spinning, etc.—processes which inelastically change the shape of a metal object, resulting in a change in its properties.  Here, I believe that the effect is more likely one of stress-relieving, in which, with the spring under a steady relatively large stress, a small number of grain boundaries which are overstressed, either break or otherwise readjust to share their stress with adjoining grains.  Over time, all grain boundaries arrive at a state which is essentially stable, but which leaves the gross material properties unchanged.  Settling noise is also generated in the joints of the instrument’s structure as they adjust to a stable state.

### 3.4  Commonly held erroneous thinking

A decade of careful research on the author's part has shown that mesoanelastic complexity disallows the creation of an ideal spring, no matter whether LeCoste zero-length, W/S leaf, or any other type. I quite agree, though the mesoanelastic effects are rather small and I claim, insignificant in practice.  Practically speaking, the mesodynamic range where we want to study earth motions is one in which the assumed parabolic (harmonic) potential well is not valid.  The assumption here seems to be that the spring motions somehow correspond to the earth motions.  In a feedback instrument any motion of the spring will be much smaller than the earth motion.  There is fine structure (FS) superposed on the harmonic potential (whether conventional or electronically ``softened''), and this FS is not even static (God help us!). Diffusion processes influenced by temperature cause the ``corrugations'' of the potential well to undergo continuous change. Bottom line-the system is not linear in the regime where we want to conduct experiments. To believe that the impulse response of the seismometer at larger levels gives a true measure of the low level sensitivity of the instrument is sheer folly. Every properly trained engineer knows that Green's function techniques do not apply to nonlinear systems.  However, for systems in which the nonlinearities are small (which includes most real-world feedback systems), linear analysis will provide useful results.  Properly trained engineers also have other techniques with which to analyze more severely non-linear systems, whenever their use becomes necessary.

### 3.5  Dithering and Stochastic resonance

It is possible that the conventional instrument may even be superior to the force-balance instrument at low levels and long periods. Why? Because of features of the new science of stochastic resonance. Engineers have long known how to deal with friction hindered systems. They use dithering, in which the instrument is shaken so as to avoid metastabilities that derive from static friction being greater than kinetic friction. Everyone knows this to be true of Coulombic friction (surfaces sliding relative to one another). It is also true, apparently, of internal friction borne of granularity (due to grain boundaries). Thus it appears that there is an advantage to operating away from critical damping, which tends to insure ``latching'' to a metastable state. Force-balancing might even be worse in forcing the system into such a state from which it does not exit.  This implies a lack of understanding as to how feedback works, probably related to the mistaken concept that feedback is somehow acting to ‘soften’ the spring.  All this discussion would seem to be appropriate to springs which undergo significant motion.  In the typical feedback instrument the movement of the spring approaches zero, and is possibly significant only at the highest frequencies.

It should be noted that latching, as  the word is used here, does not mean the instrument has "seized-up" and has no low level response at any frequency.  Rather, it means that the system refuses to oscillate at its natural frequency,  f, and evidently at other frequencies in an unknown region below f (and perhaps also above).  The system still responds to ultra-low frequencies, as evidenced by the seismometer's apparent response to tidal forcing with a 12-hr period.

## 4  Proposal

The assumption of system linearity results in a specific functional form for the free-period decay of an instrument. Specifically, linear theory requires that the logarithmic decrement be proportional to the period of oscillation raised to the first power. All long-period mechanical oscillators studied by the author deviate radically from this assumption. In particular, the log-decrement is nearly proportional to the square of the period-because of internal friction borne of anelasticity. Thus it is suggested that users of the modern instruments do experiments in which the damping is measured as a function of adjustable period.

In the STS-2, the damping at the lower corner frequenty (1/120 Hz) is controlled by the value of the proportional feedback resistor, which is normally selected for 0.707 of critical damping.  In that instrument, at 1/120Hz, the feedback electronics dictate 99.97% of the instrument response, with the spring characteristics only contributing 0.03%.  Any attempt to observe nonlinear effects by examining the overall instrument damping would be measuring the linearity of the feedback electronics, not the spring.

And the experiment would pretty much be useless, since whenever I changed the feedback elements to alter the resonant frequency, I would also have to alter the proportional feedback resistor to achieve the specified instrument damping.

In discussing the possibility of performing such an experiment with Dr. Peters, he stated that the measurements he was proposing would have to be done with the feedback turned off.

“To test my claim it is therefore necessary that you look at free-decays (exponential to first approximation) with the feedback loop disabled. “

In other words, in an operational feedback seismograph this effect of spring anelasticity can not be observed.

# Footnotes

Here Peters characterizes the spring by its Compliance,  dx/df  rather than the more commonly seen Force Constant, K df/dx.  A “soft” spring has a large  and small K.